# M1988A – Running with NI TestStand

This note is to assist customer to run M9188A with NI TestStand.

# Pre-requisite:

| Software installed | KtM9188-x64_0.4.0.0.msi (for 64-bit OS) KtM9188-x86_0.4.0.0.msi |
| --- | --- |
| DLL involved | Kyt_m9188_teststand.dll (build with .NET 4.0, suitable for NI TestStand 2013) |

# Examples:

| OutputStandardWaveform.seq | Output standard sine wave |
| --- | --- |
| OutputGroupStandardWaveform.seq | Output group standard sinewave |
| OutputArbWaveform.seq | Output arbitrary waveform and change sampling period |

# Functions:

### M9188_initialize

Initialize M9188A

| Input | resource | string | Address of the module, example: "PXI17::8::0::INSTR" |
| --- | --- | --- | --- |
|  | simulate | boolean | True = run simulation mode False = run without simulation |

### M9188_close

Close M9188 connection

### M9188_confSingleStdWaveform

Configure standard waveform for single channel

| Input | Channel | integer | Select the channel: 1 to 16 |
| --- | --- | --- | --- |
|  | OutputMode | string | Voltage = output voltage Current = output current |
|  | OutputFunction | string | DC = DC signal, when DC is selected, other parameters will be neglected Sine = Sine wave Square = Square wave Triangle = Triangle wave |
|  | Amplitude | double | Amplitude of the waveform (V or A) |
|  | OffsetLevel | double | Offset of the waveform (V or A) |
|  | PhaseShift | double | Phase shift in degree |
|  | Frequency | double | Frequency of the waveform |
|  | DutyCycleHigh | double | Duty cycle of the high level in percentage |

### M9188_setSingleOutput

Turn on or turn off the output

| Input | Channel | integer | Select the channel: 1 to 16 |
| --- | --- | --- | --- |
|  | OutputState | boolean | True = turn on False = turn off |

### M9188_confArbWaveformFromFile

Configure single channel waveform

| Input | Channel | Integer | Select the channel: 1 to 16 |
| --- | --- | --- | --- |
|  | OutputMode | string | Voltage = output voltage Current = output current |
|  | BurstCount | integer | Number of cycles to be output per triggered: 1 - 10000 |
|  | FilePath | string | Complete path of file and file name |
|  | NumOfPoints | integer | Number of data points in the data array |
|  | PlaybackMode | string | Cont = continous waveform Burst = burst waveform If this is set to Cont, then BurstCount will be ignored |
|  | SamplingPeriod | integer | Sampling period specified in us: 2us – 512us |

### M9188_changeSamplingPeriod

Changing sampling period of an arbitrary waveform

| Input | Channel | integer | Select the channel: 1 to 16 |
| --- | --- | --- | --- |
|  | SamplingPeriod | integer | Sampling period specified in us: 2us – 512us |

### M9188_confGroupStdWaveform

Configure standard waveform for a group of channels

| Input | ChannelGroup | integer | 1 = all channels ( 1 to 16) 2 = channel 1 – channel 4 3 = channel 5 – channel 8 4 = channel 9 – channel 12 5 = channel 13 – channel 16 |
| --- | --- | --- | --- |
|  | OutputMode | string | Voltage = output voltage Current = output current |
|  | OutputFunction | string | DC = DC signal, when DC is selected, other parameters will be neglected Sine = Sine wave Square = Square wave Triangle = Triangle wave |
|  | Amplitude | Double array | Amplitude of the waveform for individual channels(V or A), enable for more channels |
|  | OffsetLevel | Double array | Offset of the waveform for individual channels(V or A), enable for more channels |
|  | PhaseShift | Double array | Phase shift in degree for individual channels, enable more channels |
|  | Frequency | Double array | Frequency of the waveform for individual channels, enable for more channels |
|  | DutyCycleHigh | Double array | Duty cycle of the high level in percentage for individual channels, enable for more channels |

### M9188_setGroupOutput

Turn on or turn off the output

| Input | ChannelGroup | integer | 1 = all channels ( 1 to 16) 2 = channel 1 – channel 4 3 = channel 5 – channel 8 4 = channel 9 – channel 12 5 = channel 13 – channel 16 |
| --- | --- | --- | --- |
|  | OutputState | Boolean array | True = turn on False = turn off enable for more channels |

### M9188_setTrigger

Set trigger mode

| Input | Source | string | Set the source of the trigger Software = software trigger EXT = external hardware trigger PXI<n> = PXI lines, substitute “n” with 0-7, it can be express as “PXI2 + PXI4 + PXI5” PXI Star = Any PXI lines |
| --- | --- | --- | --- |
|  | Slope | string | Positive = positive slope Negative = negative slope |

### M9188_initiateTrigger

Initiate trigger

### M9188_setSync

Synchronize different modules of M9188A

| Input | SyncMode | string | Set as master or slave Master = master sync mode Slave = slave sync mode |
| --- | --- | --- | --- |
|  | Source | string | Set the source of the sync OFF = no sample clock EXT = external clock PXI<n> = PXI lines, substitute “n” with 0-7 |
|  | State | boolean | True = turn on False = turn off |

### M9188_setDigitalOutput

Set the 4 digital output states

| Input | BitStates | Integer array | 1 for on 0 for off Eenable the array for 4 bits |
| --- | --- | --- | --- |

### M9188_getDigitalOutput

Get the 4 digital output states

| Output | BitStates | Integer array | 1 for on 0 for off Eenable the array for 4 bits |
| --- | --- | --- | --- |

### M9188_setClockSource

Set the 10MHz system clock

| Input | Source | string | Internal = internal 10MHz oscillator External_IN = configure front clock input External_OUT = configure to output internal 10MHz to front clock PXI = PXI backplane |
| --- | --- | --- | --- |

### M9188_setEvent

Set hardware or PXI bus trigger output pulse after output per settings

| Input | Destination | string | OFF = no output trigger, same as software trigger EXT = external output trigger PXI<n> = PXI lines trigger pulse, substitute “n” with 0-7 |
| --- | --- | --- | --- |

# Flow chart