INFORMATION  ONLY 

Page 2 of 3

54832B-09A


54832B-09A 

S    E    R    V    I    C    E                N    O    T    E   







Supersedes:
NONE  

54832B
Serial Numbers:  MY00000000-MY99999999


          SG00000000-SG999999999


When Infiniium Hard Drive needs replaced for any reason. This procedure describes the process to setup the replacement hard drive in the Infiniium product. The 54855-68704 contains a Hard Drive, Infiniium software setup software and procedure. 
Disregard previous Service Notes and Service Manual part numbers that reference Hard Drive replacement part numbers.
Parts Required: 

P/N
Description
Qty.

54855-68704             Hard Drive Replacement Kit  1
Situation:

The Infiniium family has a much diversified set of hardware configurations in the installed base for each product. If a hard Drive fails or needs to be replaced for any reason then it is necessary to perform the following steps.

Solution/Action:

Setting Up the Infiniium Oscilloscope

Hard Disk Drive after Replacement

Kit Contents

• New hard disk drive with licensed Windows XP image

• Infiniium Application CD-ROM

• MakeImage CD-ROM

• Norton Ghost CD-ROM (Do not install)

• This instruction sheet
NOTE: Do not install Ghost. The Ghost software needed for making the

recovery partition is already embedded on the MakeImage CD. The Ghost CD is

included in the kit for licensing purposes only.
NOTE: The Windows XP image on the hard disk drive is licensed for use only in

Infiniium oscilloscopes and will not work in any other computer.
Replacement Procedure

1 Remove the old hard disk drive and install the new hard disk drive per the

instructions in the service guide.

2 Boot the Infiniium oscilloscope and let Windows XP find the

oscilloscope-specific hardware and load the necessary drivers.

3 Install the latest Infiniium application software from the Infiniium

Application CD-ROM.

4 Install any additional options and application licenses.

5 Perform the self test. (See the oscilloscope’s Utilities menu.)

6 Perform the self-calibration. (See the oscilloscope’s Utilities menu.)

The preceding steps made a working image on the hard disk drive. The next

step is to create a recovery image on a hidden partition.
 Creating the Recovery Partition and the Recovery Image

To make the recovery partition, you will use the MakeImage CD-ROM. The

CD-ROM will autorun and direct you to perform a Disk Cleanup so the working

partition will fit on the recovery partition.

1 Insert the MakeImage CD-ROM in the CD drive and follow the on-screen

instructions to create the recovery partition.

2 Follow the on-screen instructions to perform Disk Cleanup.

3 Follow the on-screen instructions to perform the MakeImage process. This

creates the recovery image on the recovery partition.

4 Cycle power on the Infiniium oscilloscope, and select Finish when the

dialog is displayed.

5 Remove the MakeImage CD-ROM from the CD drive.

If you encounter difficulties please visit www.agilent.com and select

Contact Us (which is in the upper right corner of the web page as of this

printing) for help with your Test & Measurement questions.

Windows is a U.S. registered trademark of Microsoft Corporation.
    ADMINISTRATIVE INFORMATION 


 

	SERVICE NOTE CLASSIFICATION:
                       INFORMATION ONLY  

	AUTHOR:  MR        PRODUCT LINE: PL1A

	 ADDITIONAL INFORMATION:  If the unit has a 10Gbit Hard Drive then the replacement is covered under warranty, while the product is in support.



[image: image1.jpg]Agilent Technologies



    © AGILENT TECHNOLOGIES, INC. 2007
    PRINTED IN U.S.A.  


     January 24, 2007
Page 1 of 3

