## Slide 1: U4611A Calibration setting study1 for output optimization
Agilent technologies
Application Engineer
Atsuki Okazaki

## Slide 2: Agenda
For output optimization
  For Near End
  For 1m cable + no channel
  For 3m cable + no channel
  For 3m cable + Device channel (5inch)
  For 3m cable + Host channel (11inch)
  For 3m cable + Device channel + CLTE On
  For 3m cable + Host channel + CLTE On
  Reference Tx compliance result

## Slide 3: For output optimization basic test setup
Test Setup
BERT
U4611A
Scope
into Std A port
From Std B port
BERT sends out CP0 pattern (no jitter, SSC on)
CP0 goes into U4611A’s std A port
U4611A repeats CP0 from Std B port
Repeated CP0 goes into scope
Capture the waveform and draw Eye diagram
Tune the output cal setting and seek for best setting

## Slide 4: For Near End optimization
BERT
U4611A
Scope
into Std A port
From Std B port
800mVp-p
Without De-emphasis
CP0
SSC on (-5000ppm)
No Jitter
Ch1 – Ch3
Second order PLL 4.9MHz
500k point
Eye height and width
U7242A +
Short cable

## Slide 5: For Near End optimization
Best setting
  Long time Constant Pre-emphasis
    Level 3 / Decay 0
  Short time Constant Pre-emphasis
    Level 0 / Decay 0
  Output level
    994mV

## Slide 6: For 3m cable + no channel optimization
BERT
U4611A
Scope
into Std A port
From Std B port
800mVp-p
Without De-emphasis
CP0
SSC on (-5000ppm)
No Jitter
Ch1 – Ch3
Second order PLL 4.9MHz
500k point
Eye height and width
U7242A +
3m cable

## Slide 7: For 3m cable + no channel optimization
Best setting
  Long time Constant Pre-emphasis
    Level 11 / Decay 1
  Short time Constant Pre-emphasis
    Level 0 / Decay 0
  Output level
    582mV

## Slide 8: For 3m cable + Device channel (5inch)
BERT
U4611A
Scope
into Std A port
800mVp-p
Without De-emphasis
CP0
SSC on (-5000ppm)
No Jitter
Ch1 – Ch3
Second order PLL 4.9MHz
500k point
Eye height and width
U7242A + short cable
Infiniisim
S-para simulation
De-embed “U7242+short cable”
& Embed “3m+device channel”

## Slide 9: For 3m cable + device channel optimization
Best setting
  Long time Constant Pre-emphasis
    Level 0 / Decay 0
  Short time Constant Pre-emphasis
    Level 14 / Decay 6
  Output level
    216mV

## Slide 10: For 3m cable + Host channel (11inch)
BERT
U4611A
Scope
into Std A port
800mVp-p
Without De-emphasis
CP0
SSC on (-5000ppm)
No Jitter
Ch1 – Ch3
Second order PLL 4.9MHz
500k point
Eye height and width
U7242A + short cable
Infiniisim
S-para simulation
De-embed “U7242+short cable”
& Embed “3m+Host channel”

## Slide 11: For 3m cable + host channel optimization
Best setting
  Long time Constant Pre-emphasis
    Level 0 / Decay 0
  Short time Constant Pre-emphasis
    Level 14 / Decay 6
  Output level
    216mV

## Slide 12: For 3m cable + Device channel + CLTE
BERT
U4611A
Scope
into Std A port
800mVp-p
Without De-emphasis
CP0
SSC on (-5000ppm)
No Jitter
Ch1 – Ch3
Second order PLL 4.9MHz
500k point
Eye height and width
U7242A + short cable
Infiniisim
S-para simulation
De-embed “U7242+short cable”
& Embed “3m+device channel”
CTLE
SIM

## Slide 13: For 3m cable + device channel + CLTE optimization
Best setting
  Long time Constant Pre-emphasis
    Level 3 / Decay 2
  Short time Constant Pre-emphasis
    Level 0 / Decay 0
  Output level
    994mV

## Slide 14: For 3m cable + Host channel + CLTE
BERT
U4611A
Scope
into Std A port
800mVp-p
Without De-emphasis
CP0
SSC on (-5000ppm)
No Jitter
Ch1 – Ch3
Second order PLL 4.9MHz
500k point
Eye height and width
U7242A + short cable
Infiniisim
S-para simulation
De-embed “U7242+short cable”
& Embed “3m+host channel”
CTLE
SIM

## Slide 15: For 3m cable + host channel + CLTE optimization
Best setting
  Long time Constant Pre-emphasis
    Level 6 / Decay 0
  Short time Constant Pre-emphasis
    Level 0 / Decay 0
  Output level
    808mV

## Slide 16: Result
 | Long Level | Long Decay | Short Level | Short Decay | Output Level
Near End | 3 | 0 | 0 | 0 | 994
3m cable | 11 | 1 | 0 | 0 | 582
3m + Device | 0 | 0 | 14 | 6 | 216
3m + host | 0 | 0 | 14 | 6 | 216
3m + Device + CLTE | 3 | 2 | 0 | 0 | 994
3m + Host + CLTE | 6 | 0 | 0 | 0 | 892

## Slide 17: TX Compliance
 | Long Level | Long Decay | Short Level | Short Decay | Output Level
3m + Device + CLTE | 3 | 2 | 0 | 0 | 994
All Pass