This guide is divided in 4 Chapters, please follow the described order.

- De-Install Aurora driver

- Install USB Device Driver

- Upgrade G4301A Firmware/ Install shortcut for new SFC Utility SW

- Upgrade G4303A SFC ALS Firmware

- Install/Upgrade OL CDS C.01.05 with HotFix 2 or higher and driver pack A.02.08

Necessary software and firmware files can be downloaded from the Support Portal and Field Server except the OL CDS A.01.05 which can be downloaded from Subscribe Net and HotFix 2 or higher for OL CDS A.01.05, which can be found via Service Note Viewer,

(Service Note SN OPENLAB CDS ChemStation EDITION-024)

- De-Install ‘SFC Fusion Chemstation Interface’ via add/remove programs or uninstall software (depending on the OS of the PC)

- New UsbDevice Driver

New Usb device Driver necessary to support the Firmware 0307:

Agilent.Sfc.UsbDeviceDriver_x64.msi

  - For Controlling PC OS Win7 64bit

Agilent.Sfc.UsbDeviceDriver_x86.msi

  - For Controlling OS with any other OS as Win7 64bit

Caution: The files ‘infinity1260sfc’ and ‘rootsupd’ need to be present in the folder ‘UsbDeviceDriver’ to install the device driver

- Upgrade Firmware for G4301A SFC Control Module

Copy folder ‘FirmwareUpdateTool’ on the HDD of the controlling PC and create Shortcut on the Desktop for ‘Agilent.Sfc.FirmwareUpdateTool’

Use shortcut ‘Agilent.Sfc.FirmwareUpdateTool’ to open FW update tool and upgrade FW

- 4301A_B2_A307_003.afi  for Non-Evolution Modules/ Mainboards rev B

- 4301A_C2_A307_003.afi  for Evolution Modules/ Mainboards rev C

Caution: New FW A0307 is not backwards compatible, works only in the new driver environment

Copy folder ‘TestUtility’ on the HDD of the controlling PC and create Shortcut on the Desktop for ‘Agilent.NonLCDrivers.Sfc.Utility’

Use shortcut to open ‚ Agilent.NonLCDrivers.Sfc.Utility’ if you need to perform pressure sensor calibration or G4301A performance tests. (Not necessary to do calibration after SW and FW upgrade only.)

- New Firmware necessary for G4303A SFC Autosampler

New FW ‚4303A_A654_003.dlb‘ is mandatory otherwise G4303A appears offline in OL CDS after driver installation A.02.08.

Use Lab-Advisor to upgrade G4303A FW

- Install/Upgrade OL CDS C.01.05[35] with HotFix 2[37]  and driver pack A.02.08

Make sure that OL CDS A.01.05 [37] or higher is installed on the controlling PC

before installing the driver A.02.08 or higher.

Caution, with driver A.02.08:

- No support for B.xx.xx Chemstation, upgrade to C.01.05 necessary

- Only RC.Net, no Classic Chemstation

Install driver A.02.08 or higher from folder ‘READY_LCINSTALLPACKAGE_02.08.424’

By double clicking ‘Agilent_OpenLAB_CDS_ChemStation_LC_and_CE_Drivers.msi’

Configure the system with Auto-configuration, all modules in the system should be found except the G4301A SFC Control module. This module has to be manually moved in the configuration and the COM Port need to be chosen.

(If you are not sure which COM Port is used by the G4301A open SFC Test utility SW, here you will see in the home screen the right Com port which should be chosen).

Note: 	Methods created on older SFC driver versions are not compatible with the new SFC Driver, these ‘old’ methods need to be loaded after the driver upgrade, the G4303A (injector program not necessary anymore), G4302A (solvent compressibility) and G4301A settings need to be reviewed and corrected if necessary, then the methods need to be saved.