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12000A Operation Manual Addendum

Pulse Modulation Operation

Description:
Output exhibits a level drop when changing power levels while pulse modulation is active. Pulse level is slow to level for low duty cycles

Firmware Version:
V0.72 B47-26 to current (V0.72 B49-19). 

Correction:
For firmware versions listed above, proper pulse modulation operation is obtained by setting the RF Retrace to the ON position. The RF Retrace setting is located in the Configuration Menu. 

From any menu,

Press the CONFIG button on the front panel.

Using the cursor controls, move the cursor to the RF Retrace setting.

Using the rotary knob or arrow keys, switch the RF Retrace to the ON position.

Exit the configuration menu and continue normal operation.

