November 16, 2001


12000A Operation Manual Addendum

Pulse Modulation Power Level Change

Description:
Pulse modulation disables when changing power level setting, LED indicator remains active. 

Firmware Version:
V0.72 B49-19

Correction:
Power level can be adjusted while pulse modulation is active. The level must be adjusted from the Fixed Frequency menu only. Adjusting level from the Fixed Power menu will disable pulse modulation. Changing level in the Fixed Frequency menu will reactivate pulse modulation. The following operation is required in order to change power level while pulse modulation is active:

Press the CW button

Using the cursor keys, move the cursor over the Power Level setting

Adjust power to the desired output level.

